## SHA Table
SHA ID Number | Hazardous Situation ID | Hazard | Hazardous Situation | Function | OTS Software (if OTS, identify component) | Hazardous sequence of events | S/W Related Cause(s) | Harm | Severity | Exploitability - (Cyber) (Pre-Mitigation) | Probability of Harm (software/Use-Related) (Pre-Mitigation) | Initial Risk Rating | Risk Control Measures | Demonstration of Effectiveness (Trace to Verification) | Severity of Harm (Post-Mitigation) | Exploitability - (Cyber) | Probability of Harm (software/Use-Related) | Final Risk Rating | New HS if applicable If yes, reference new row with SHA ID | System DFMEA Trace | SRA Link | URRA Item | Residual Risk Acceptability (Rationale for Acceptability per GQP-10-02, Risk Management Report)
HAZ-PUMP-101 | HS-PUMP-101 | Unintended bolus delivered while pump is in standby. | Pump delivers an armed PCA bolus while in standby, unattended by the clinician. | Bolus delivery control | None | 1. Clinician arms a PCA bolus. 2. Pump transitions to standby before delivery. 3. Standby transition fails to clear the armed-bolus flag. 4. The armed bolus is delivered unattended. | Standby state-transition routine does not reset the armed-bolus flag (state-machine defect). | Opioid overdose; respiratory depression. | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-101: the standby transition SHALL clear any armed bolus and SHALL record the clear event. | TC-PUMP-101-A (armed bolus cleared on standby), TC-PUMP-101-B (clear event logged), TC-PUMP-101-C (boundary: standby entered during the arming window). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-11 | N/A - no cyber cause | URRA-PUMP-01 | Acceptable per GQP-10-02; verified clearing of the armed-bolus flag on every standby transition reduces probability to Remote; residual risk is outweighed by the clinical benefit of PCA therapy.
HAZ-PUMP-102 | HS-PUMP-102 | Downstream occlusion not detected; analgesia silently interrupted. | Occlusion goes undetected and PCA analgesia is silently interrupted. | Occlusion detection | None | 1. Downstream line occludes. 2. Line pressure rises above 15 psi. 3. The occlusion detector must alarm within 2 seconds. 4. If undetected, prescribed analgesia is not delivered. | Pressure-sampling task could miss the 2-second alarm deadline under peak CPU load. | Untreated pain; inadequate analgesia. | Serious | N/A - no network exposure | Occasional | Medium | REQ-PUMP-102: an occlusion SHALL be detected and alarmed within 2 seconds of line pressure exceeding 15 psi. | TC-PUMP-102-A (alarm within 2 s), TC-PUMP-102-B (boundary: 15.0 vs 15.1 psi), TC-PUMP-102-C (negative: sensor disconnect raises a fault). | Serious | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-12 | N/A - no cyber cause | URRA-PUMP-02 | Acceptable; 2-second bounded detection verified under worst-case scheduling load reduces probability to Remote.
HAZ-PUMP-103 | HS-PUMP-103 | Dose computed from an out-of-range drug concentration. | Pump computes and delivers a dose from an out-of-range drug concentration. | Drug concentration entry and dose calculation | None | 1. Clinician enters the drug concentration. 2. A value outside the approved 0.1-10 mg/mL range is entered. 3. Dose calculation uses the out-of-range value. 4. The pump delivers an incorrect dose. | Missing range validation on the commanded concentration input allows an out-of-range value into the dose calculation. | Overdose or underdose of analgesic. | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-103: the pump SHALL reject drug concentrations outside 0.1-10 mg/mL and SHALL require re-entry. | TC-PUMP-103-A (valid value accepted), TC-PUMP-103-B (boundary: 0.1 and 10.0 accepted; 0.09 and 10.1 rejected), TC-PUMP-103-C (negative: non-numeric input rejected). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-13 | N/A - no cyber cause | URRA-PUMP-03 | Acceptable; enforced range validation blocks out-of-range concentrations, reducing probability to Remote.
HAZ-PUMP-104 | HS-PUMP-104 | Cumulative in-line air exceeds the safe limit without a pump halt. | Cumulative infused air exceeds the 100 microliter safe limit without a pump halt. | Air-in-line detection | None | 1. Air enters the infusion line. 2. The air-volume accumulator tracks cumulative air. 3. The pump must halt when cumulative air exceeds 100 microliters. 4. If the accumulator wraps and under-counts, air is infused. | The air-volume accumulator could overflow and wrap, under-counting cumulative air. | Venous air embolism. | Catastrophic | N/A - no network exposure | Remote | High | REQ-PUMP-104: the pump SHALL halt infusion and alarm when cumulative in-line air exceeds 100 microliters; the accumulator SHALL saturate rather than wrap. | TC-PUMP-104-A (halt above 100 uL), TC-PUMP-104-B (boundary: 100 vs 101 uL), TC-PUMP-104-C (negative: accumulator saturates at max, no wrap). | Catastrophic | N/A - no network exposure | Improbable | Low | None | FMEA-PUMP-14 | N/A - no cyber cause | URRA-PUMP-04 | Acceptable; saturating accumulator and bounded halt verified, reducing probability to Improbable.
HAZ-PUMP-105 | HS-PUMP-105 | Uncontrolled gravity free-flow when the cassette is removed. | Uncontrolled gravity free-flow of drug occurs when the cassette is removed. | Free-flow protection | None | 1. Clinician opens the pump door. 2. The anti-free-flow valve must engage before the latch releases. 3. Software must confirm the valve closed before releasing the door latch. 4. If the latch releases first, gravity free-flow delivers uncontrolled drug. | Door-latch release could be commanded before the valve-closed confirmation is read. | Overdose. | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-105: the door latch SHALL NOT release until the anti-free-flow valve reports closed. | TC-PUMP-105-A (latch blocked until valve closed), TC-PUMP-105-B (boundary: valve-close confirmation timeout), TC-PUMP-105-C (negative: valve fault holds the latch closed). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-15 | N/A - no cyber cause | URRA-PUMP-05 | Acceptable; interlock verified so the latch cannot release before valve confirmation, reducing probability to Remote.
HAZ-PUMP-106 | HS-PUMP-106 | Pump stops on battery depletion without a prior alarm. | Pump stops on battery depletion without raising a prior low-battery alarm. | Power management and low-battery handling | None | 1. Mains power is lost. 2. The battery discharges. 3. At 10% state-of-charge the pump must alarm and continue for at least 30 minutes. 4. If it stops silently, therapy is interrupted unnoticed. | The battery gauge could report a stale state-of-charge if its sampling task starves. | Untreated pain; therapy interruption. | Serious | N/A - no network exposure | Occasional | Medium | REQ-PUMP-106: at 10% state-of-charge the pump SHALL raise a low-battery alarm and SHALL sustain infusion for at least 30 minutes. | TC-PUMP-106-A (alarm at 10%), TC-PUMP-106-B (boundary: 10 vs 9%), TC-PUMP-106-C (negative: gauge fault triggers a conservative alarm). | Serious | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-16 | N/A - no cyber cause | URRA-PUMP-06 | Acceptable; bounded alarm and 30-minute sustain verified, reducing probability of silent interruption to Remote.
HAZ-PUMP-107 | HS-PUMP-107 | Demand bolus delivered within the programmed lockout interval. | A demand bolus is delivered within the programmed lockout interval, stacking doses. | PCA lockout enforcement | None | 1. Patient presses the demand button and a bolus is delivered. 2. Within the lockout interval the patient presses again. 3. The lockout must ignore the second request. 4. If the lockout is bypassed, a second bolus stacks. | The lockout timer could be reset by a spurious button-debounce event, re-enabling an early demand. | Opioid overdose. | Critical | N/A - no network exposure | Probable | High | REQ-PUMP-107: demand boluses within the programmed lockout interval SHALL be ignored and logged; debounce SHALL NOT reset the lockout timer. | TC-PUMP-107-A (early demand ignored), TC-PUMP-107-B (boundary: request at exactly lockout expiry accepted; 1 s before rejected), TC-PUMP-107-C (negative: rapid double-press does not reset the timer). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-17 | N/A - no cyber cause | URRA-PUMP-07 | Acceptable; lockout enforcement verified against debounce interference, reducing probability of dose stacking to Remote.
HAZ-PUMP-108 | HS-PUMP-108 | Programmed rate retained instead of KVO after infusion completion. | Pump continues at the programmed rate after infusion completion instead of the keep-vein-open rate. | Post-infusion keep-vein-open (KVO) control | None | 1. The programmed infusion completes. 2. The pump transitions to the KVO rate. 3. The KVO setpoint (<=5 mL/hr) must be loaded. 4. If the prior programmed rate is retained, over-infusion continues. | The KVO transition could retain the prior programmed rate instead of loading the KVO setpoint. | Over-infusion; overdose. | Serious | N/A - no network exposure | Occasional | Medium | REQ-PUMP-108: on infusion completion the pump SHALL load the KVO setpoint (<=5 mL/hr) and SHALL NOT retain the programmed rate. | TC-PUMP-108-A (KVO setpoint loaded on completion), TC-PUMP-108-B (boundary: 5.0 vs 5.1 mL/hr), TC-PUMP-108-C (negative: setpoint-load failure alarms and stops). | Serious | N/A - no network exposure | Remote | Low |  | FMEA-PUMP-18 | N/A - no cyber cause | URRA-PUMP-08 | Acceptable; verified KVO setpoint load on completion reduces over-infusion probability to Remote. NOTE: a distinct newly-observed situation - KVO transition while an alarm-silence is active - was raised in design review but has not yet been captured as its own SHA row.
HAZ-PUMP-109 | HS-PUMP-109 | Drug leaks from a cracked cassette housing, under-delivering the dose. | Drug leaks from a cracked cassette housing, and the prescribed dose is not fully delivered. | Fluid path containment (mechanical) | None | 1. The cassette housing develops a hairline crack under mechanical stress. 2. Drug leaks from the fluid path. 3. The prescribed dose is not fully delivered to the patient. | None - purely mechanical cause; no software contribution. | Under-dose; inadequate analgesia. | Serious | N/A - no software or cyber cause | Remote | Medium | REQ-PUMP-109: the cassette housing SHALL withstand three times maximum operating pressure without leakage per mechanical qualification, and incoming inspection SHALL verify housing integrity. | MECH-TC-109-A (pressure burst test), MECH-TC-109-B (drop test), incoming-inspection AQL sampling. | Serious | N/A - no software or cyber cause | Improbable | Low | None | DFMEA-PUMP-M07 | N/A - no cyber cause | URRA-PUMP-09 | Acceptable; mechanical qualification and incoming inspection reduce leakage probability to Improbable. Software verification is not the closure mechanism because the cause is mechanical.
HAZ-PUMP-110 | HS-PUMP-110 | Luer connector separates and analgesia is not delivered. | The luer connector separates under vibration and analgesia is not delivered to the patient. | Fluid path connection integrity (mechanical) | None | 1. Transport vibration loosens the luer connector. 2. The connection separates. 3. Drug is not delivered to the patient. | None - mechanical connector loosening; no software contribution. | Under-dose; untreated pain. | Serious | N/A - no software or cyber cause | Remote | Medium | REQ-PUMP-110: the luer connection SHALL be a locking type qualified to retain integrity under IEC 60068 vibration, and labeling SHALL instruct connection verification. | MECH-TC-110-A (IEC 60068 vibration retention), MECH-TC-110-B (locking-luer pull-off force), labeling verification. | Serious | N/A - no software or cyber cause | Improbable | Low | None | DFMEA-PUMP-M08 | N/A - no cyber cause | URRA-PUMP-10 | Acceptable; locking luer qualified under vibration and reinforced by labeling reduces separation probability to Improbable. Software verification is not the closure mechanism because the cause is mechanical.
HAZ-PUMP-111 | HS-PUMP-111 | Alarm not annunciated. | The alarm is not annunciated. | Alarm annunciation | None | Alarm condition occurs. | Audio driver defect. | Delayed clinician response to a hazard. | Serious | N/A - no network exposure | Occasional | Medium | REQ-PUMP-111: audible alarms SHALL reach 65 dBA at 1 metre. | TC-PUMP-111-A (alarm loudness at 1 m). | Serious | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-21 | N/A - no cyber cause | URRA-PUMP-11 | Acceptable; alarm loudness verified at 65 dBA reduces the probability of a missed alarm to Remote.
HAZ-PUMP-112 | HS-PUMP-112 | Infusion status display freezes. | The infusion status display shows stale values. | Display of infusion status | None | 1. The UI rendering thread stalls. 2. The status display stops updating. 3. Displayed values become stale. | UI thread priority inversion stalls rendering. | TBD | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-112: the infusion status display SHALL refresh at least once per second. | TC-PUMP-112-A (display refresh rate), TC-PUMP-112-B (negative: rendering-thread watchdog recovers the UI). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-22 | N/A - no cyber cause | URRA-PUMP-12 | Acceptable; a rendering watchdog recovers a stalled UI and bounds staleness, reducing probability to Remote.
HAZ-PUMP-113 | HS-PUMP-113 | Incorrect infusion rate commanded. | Pump runs at an incorrect commanded rate. | Dose rate command | None | 1. The rate setpoint is computed. 2. A software fault corrupts the commanded rate. 3. The pump runs at the wrong rate. 4. Over- or under-infusion results. | Software error. | Over- or under-infusion. | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-113: the commanded rate SHALL be range-checked and cross-verified by an independent monitor. | TC-PUMP-113-A (range check), TC-PUMP-113-B (boundary: max rate), TC-PUMP-113-C (negative: monitor mismatch stops the pump). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-23 | N/A - no cyber cause | URRA-PUMP-13 | Acceptable; independent rate monitor and range check reduce the probability of an undetected wrong rate to Remote.
HAZ-PUMP-114 | HS-PUMP-114 | Unauthorized modification of dose limits. | Pump operates with an unsafe, unauthorized dose limit. | Wireless dose-limit configuration (network-connected) | None | 1. The pump receives a dose-limit configuration over the hospital network. 2. A malformed, unauthenticated message alters the hard limit. 3. The pump accepts an unsafe dose limit. 4. Overdose occurs on subsequent programming. | Parsing bug in the configuration message handler. | Overdose. | Critical | High | Occasional | High | REQ-PUMP-114: configuration messages SHALL be schema-validated before any limit is applied. | TC-PUMP-114-A (malformed message rejected), TC-PUMP-114-B (boundary: out-of-range limit rejected). | Critical | Medium | Remote | Medium | None | FMEA-PUMP-24 | SRA-PUMP-14 | URRA-PUMP-14 | Acceptable; schema validation of configuration messages reduces the probability of an accepted malformed limit to Remote.
HAZ-PUMP-115 | HS-PUMP-115 | Occlusion undetected due to a threshold comparison error. | An occlusion goes undetected because the pressure threshold comparison is wrong. | Occlusion pressure limit | None | 1. Downstream pressure rises. 2. An off-by-one in the threshold comparison suppresses the alarm at the limit. 3. The occlusion is not detected. 4. Analgesia is interrupted. | Pressure-threshold comparison off-by-one at the 15 psi limit. | Untreated pain; inadequate analgesia. | Catastrophic | N/A - no network exposure | Frequent | Acceptable | REQ-PUMP-115: the occlusion threshold comparison SHALL use an inclusive limit and be unit-tested at the boundary. | TC-PUMP-115-A (boundary at 15 psi), TC-PUMP-115-B (alarm fires above limit). | Serious | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-25 | N/A - no cyber cause | URRA-PUMP-15 | Acceptable; corrected inclusive comparison verified at the boundary reduces the probability of a missed occlusion to Remote.
HAZ-PUMP-116 | HS-PUMP-116 | Spoofed telemetry masks an active alarm state. | A clinician relies on spoofed telemetry that hides an active alarm. | Remote monitoring telemetry (network-connected) | None | 1. The pump publishes alarm/telemetry state to a remote monitor. 2. An attacker injects spoofed telemetry because message authenticity is not enforced. 3. The remote monitor shows a normal state. 4. The clinician's response to a real alarm is delayed. | Telemetry message authenticity is not enforced (cybersecurity cause). | Delayed response to a hazard. | Serious | High | Occasional | Medium | REQ-PUMP-116: telemetry messages SHALL be authenticated and integrity-protected. | TC-PUMP-116-A (spoofed telemetry rejected), TC-PUMP-116-B (integrity failure raises a fault). | Serious | Medium | Remote | Low |  | FMEA-PUMP-26 |  | URRA-PUMP-16 | Acceptable; authenticated telemetry reduces the probability of an accepted spoofed state to Remote. NOTE: the same missing message-authenticity control plausibly implies a distinct newly-identified hazardous situation - acceptance of spoofed control commands, not merely spoofed telemetry - which has not yet been captured as its own SHA row.
HAZ-PUMP-117 | HS-PUMP-117 | Pump commits an unintended dose value. | The pump commits and delivers an unintended dose value. | Keypad dose entry | None | 1. Clinician types the dose. 2. Keypad debounce may drop a digit. 3. The committed dose is not independently cross-checked before delivery. 4. An unintended dose is delivered. | (a) Keypad debounce may drop a keypress; (b) the committed dose is not independently cross-checked before delivery. | Under- or over-dose. | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-117: keypad input SHALL apply a 20 ms debounce filter. | TC-PUMP-117-A (debounce filtering), TC-PUMP-117-B (boundary: 20 ms chatter window). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-27 | N/A - no cyber cause | URRA-PUMP-17 | Acceptable; debounce filtering verified reduces dropped-digit probability to Remote.
HAZ-PUMP-118 | HS-PUMP-118 | Infusion context lost during a battery hot-swap. | Infusion context is lost during a battery hot-swap and delivery is interrupted or corrupted. | Battery hot-swap | None | 1. Mains power is removed and the battery is hot-swapped. 2. The power-fail handler must persist the infusion context to non-volatile memory. 3. On restore the context must be reloaded. 4. If not persisted, delivery state is corrupted. | The power-fail handler may not persist the infusion context before the supercapacitor bridge expires. | Interrupted or corrupted infusion; over- or under-dose. | Critical | N/A - no network exposure | Occasional | High | Software mitigations are in place to preserve the infusion context across a battery hot-swap. | TC-PUMP-118-A (context persisted on power-fail), TC-PUMP-118-B (boundary: bridge-time limit), TC-PUMP-118-C (negative: persistence failure alarms and stops). | Critical | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-28 | N/A - no cyber cause | URRA-PUMP-18 | Acceptable; verified context persistence within the supercapacitor bridge time reduces corruption probability to Remote.
HAZ-PUMP-119 | HS-PUMP-119 | Pump overheats without thermal shutdown. | The pump internal temperature exceeds the safe limit without a thermal shutdown. | Over-temperature shutdown | None | 1. Internal temperature rises. 2. The thermal-cutoff task must shut the pump down at 60 degrees C. 3. If it does not trigger, the pump continues to heat. 4. Drug degradation or a contact burn results. | The thermal-cutoff task may not trigger shutdown at the 60 degrees C limit. | Drug degradation; contact burn. | Serious | N/A - no network exposure | Occasional | Medium | REQ-PUMP-119: the pump SHALL shut down and alarm when internal temperature exceeds 60 degrees C. | TC-PUMP-119-A verifies the pump powers on and displays the current temperature. | Serious | N/A - no network exposure | Remote | Low | None | FMEA-PUMP-29 | N/A - no cyber cause | URRA-PUMP-19 | Acceptable; thermal cutoff bounds the maximum temperature, reducing the probability of overheating to Remote.
HAZ-PUMP-120 | HS-PUMP-120 | Hard-limit override delivers an unsafe dose. | A hard-limit override allows an unsafe dose to be programmed and delivered. | Dose-error reduction (guardrails) | None | 1. Clinician attempts a dose above the hard limit. 2. The guardrail must block the override. 3. If the override is allowed without a second check, an unsafe dose is programmed. 4. Overdose results. | The hard-limit override path may not require an independent second confirmation. | Overdose. | Critical | N/A - no network exposure | Occasional | High | REQ-PUMP-120: a hard-limit override SHALL require an independent second-clinician confirmation and SHALL be logged. | TC-PUMP-120-A (override blocked without second confirmation), TC-PUMP-120-B (boundary: at the hard limit), TC-PUMP-120-C (negative: override logged). | Critical | N/A - no network exposure | TBD | Low | None | FMEA-PUMP-30 | N/A - no cyber cause | URRA-PUMP-20 | Acceptable.